## Slide 1
Slides by Eng. Enrico Maria Randone, PhD.
Dipartimento di Elettronica - Università di Pavia
Tel. 0382 98 5224 e-mail: enrico.randone@unipv.it
It is acknowledged the origin of some figures, photos and texts
as didactic material originally produced by Prof. Silvano Donati (Università di Pavia)
and/or Dr. Eng. Guido Giuliani (Università di Pavia)
”LASER INTERFEROMETRY” part 3 - ESPI

## Slide 2: ESPI (Electronic Speckle Pattern Interferometry)
Each single speckle can be used as a SINGLE INTERFEROMETRIC CHANNEL, and this feature can be used to develop single point instruments like a Vibrometer (we must always remember the dynamic limitations due to speckles dimensions)
When a single-mode laser source shines on a diffuser, however, more than one speckle is diffused back to the observer, and the number of the speckles depends on the dimension of the laser spot on the diffuser
Is it possible to develop an instrument that is able to process all the speckles in parallel? In this way m x n (pixels) interferometric images of the movement of the diffuser could be obtained…

## Slide 3: ESPI: Measuring Principle (I)
The diffusing surface is illuminated by a laser source, and the whole speckled image is observed with an imaging system (CCD camera), after superposition with the laser reference beam
In order to develop a great number of speckles, particular optics should be used, like double-lenses telescopes: instead of focusing the laser beam on the surface, like in a single point Vibrometer, a big area of the diffuser must be illuminated
He-Ne LASER

## Slide 4: ESPI: Measuring Principle (II)
A single STATIC image from the camera doesn’t contain useful information…
But what happens if the surface starts moving or swinging back-and-forth?
As suggested before, each speckle can be used as a single interferometric channel: each illuminated point of the diffuser undergoes displacements that are shorter than the longitudinal dimensions of the speckles (NEDsp)
The idea is: light from each speckles interferes with the reference laser beam, and “inside” of each spot fringes will develop: high intensity medium intensity dark medium intensityhigh intensity  …
Each speckle will behave as a typical interferometer based on fringes!

## Slide 5: ESPI “Time-Averaging” Technique (I)
The target, typically a mechanical membrane, is actuated with a command signal. This produces, on its surface, stationary waves related to the possible vibrating mode of the membrane/material
The detected video signal is processed to get a sort of “mean of the intensity” related to each speckle, that can be considered a “pixel” of the image: in the Nodes of the surface (points at rest) the speckles intensity is constant (“bright” regions), whereas in the Anti-Nodes (points that exhibit the maximum excursion) the speckles intensity varies with time (“darker” regions) with an average intensity halved or anyway reduced
Reference Image of
a loudspeaker used
as a diffusing target (Target at rest)
Increasing Frequency
200 Hz
400 Hz
600 Hz
2000 Hz
2200 Hz

## Slide 6: ESPI “Time-Averaging” Technique (II)
The technique is extremely easy to set-up 
Vibrating Modes of a target area of hundreds of squared centimeters are directly/parallel observed 
A direct measure of the amplitude of the vibrations induced in the target is NOT achieved, as we get only the Nodes and Anti-Nodes position and distribution 
Nodes and Anti-Nodes distributions are frequently analyzed in mechanics engineering and the technique is indeed very useful 

## Slide 7: ESPI “Frame-Subtraction” Technique (I)
The idea is to compare two static images from the camera, both of which, singularly, don’t gives any information on the target surface:
  A reference image (target at rest)
  A measuring image (target subjected to a static deformation)
If we subtract pixel-to-pixel the two images, we obtain a new image that contains the interferometric intensity: variation experienced by each speckle during the movement (it must always be sm < sl)
The resulting image contains the information on the surface static deformation in terms of bi-dimensional interferometric fringes, where each fringe = /2
Reference Image (un-deformed Target)
Measuring Image (deformed-target)
Subtraction Image
(bi-dimensional fringes
Distribution)
/2

## Slide 8: ESPI “Frame-Subtraction” Technique (II)
/2
Vibration measurement
Subtraction image
Fringes analysis gives a direct idea of the deformation/vibration amplitude
Dynamical limitations due to the longitudinal speckle dimension
What do we need?
  ‘70: video-recorder + electronic “hardware” subtraction
  Today: frame-grabber + PC + subtraction “software” algorithm
This Technique can be applied also to variable target deformation: a “real-time” software approach must be implemented for video-processing

## Slide 9: ESPI “Frame-Subtraction” Technique (III)
Ambiguity on the sign of the target deformation
The problem is overcome thanks to a technique called “phase unwrapping”
Two possible implementation:
  Software
  Hardware (Optical phase-shift)
    second acquisition with a /16 displacement
    of the beam splitter used for reference/measuring
    fields in order to get a bi-dimensional distribution
    of the fringes having a phase shift of /2 (/8)
    with respect to the first acquisition
    sin(φ) & cos(φ) !!!

## Slide 10: ESPI “Frame-Subtraction” Technique (IV)
Examples of ESPI frame-subtraction images obtained from a static mechanical deformation (left) or from deformation due to a temperature gradient (right) applied to the target

## Slide 11: ESPI “Frame-Subtraction” Technique (V)
laser = 632.8 nm
Point A is always unperturbed by the deformation applied to the target (“reference-phase point”)
What are the possible displacements of points B, C, D?
A
C
B
D
A
C
B
D
A
C
B
D