## Slide 1: “Optical Measurements” Master Degree in Engineering Automation-, Electronics-, Physics-, Telecommunication- Engineering
Optical Measurement Instrumentation Prof. Cesare Svelto Politecnico di Milano
Notes: Buongiorno a tutti,
sono Cesare Svelto
del Politecnico di Milano.
Questa mattina vi parlerò delle sorgenti laser in continua e delle pricipali misure per la loro caratterizzazione.
Descriveremo la struttura dei laser a stato solido e dei laser a semoconduttore (diodi laser)
per la caratterizzazione di sorgenti laser in continua e dei campioni ottici di frequenza.
La prima parte del mio intervento sarà seguita da una lezione del Prof. Elio Bava sui modelli descrittivi dell’instabilità di frequenza e sulle relative tecniche di misura.
Nella seconda metà della mia presentazione descriverò, dal punto di vista sperimentale, alcuni metodi ed esperienze di stabilizzazione e la realizzazione di campioni ottici di frequenza.

## Slide 2: Summary
Optical Power Meter
CCD detectors (Charge Coupled Device)
and measurement of laser beam profile
Wave-Meter and spectrometer/monochromator
Optical Spectrum Analyzer (OSA)
Optical Time Domain Reflectometry (OTDR)
Meas. of Insertion Loss
Meas. of Polarization Mode Dispersion (PMD)
Meas. of Bit Error Rate (BER)
Notes: Dopo una breve introduzione sull’argomento e sulle principali caratteristiche degli oscillatori laser,
vedremo più in dettaglio la struttura di una sorgente laser, i principi fisici che permettono l’azione laser, e i risonatori ottici.
A questo punto descriveremo in dettaglio alcune proprietà caratteristiche dei fasci laser.
Passeremo quindi alle misure della potenza ottica indicando metodi e strumenti adatti allo scopo.
Di seguito vedremo come si effettua la caratterizzazione del profilo spaziale e della divergenza di un fascio laser.
Infine illustreremo metodi e strumenti per la misura della lunghezza d’onda e dello spettro di emissione.
Così si concluderà la prima parte di questo intervento.

## Slide 3: Meas. of Optical Power (DEFINITIONS and METHODS)
E Electric field [V/m]
Laser
calibrated detector
voltage
output
Vu
calibrati
attenuators
P
P
2
1
3
Vu = P
P = Power [W]
I = Intensity [W/m2]
0 = (0/0)1/2 = 377 
vacuum impedance
ADC
dig. display
for a PD it is  = (A/W) R(V/A)


h
e
  

## Slide 4: Optical Power Meter (structure)
detection
scheme
device
The detector can be THERMAL or SEMICONDUCTOR
(thermal detectors have “flat” response)
CALIBRATION
spatial filter
attenuator
commercial
device
0.1 dB = 2.310-2
0.01 dB = 2.310-3
absPopt=Ptherm= KT=K(T-Tamb)
 T  Popt
“calibrated” Si

## Slide 5: Power Meter (thermal sensor)
THERMOPILE detector
thermal profile
reading head
The thermal field has a radial gradient: from central “hot”
zone to the “cold” border
Very “slow” (B=1 Hz) but accurate for DC power or average meas.

## Slide 6: Power Meter (semiconductor specs.)
Pmin=-90dBm=1pW
Pmax=+21.5dBm=140mW 11 orders of magnitude!!!
READING HEAD
it’s a good Digital Volt Meter!
READING BODY

## Slide 7: CCD detectors
A 1024x1024 CCD camera for UV spectrometer
Sliding and collection of photodetected charges
Image reconstruction on a pixel matrix
(visualization by raster display)
Important: single pixel dimensions
Figure from book ‘Photodetectors',
by S.Donati, 2nd Ed., AEI Milan 1997.

## Slide 8: Meas. of laser beam profile ( Analysis in the xy plane during z-propagation )
Meas. of the beam spot size
Far field (z>>zR): meas. of divergence
Meas. of astigmatism w0x  w0y x  y
Meas. of M2 =

## Slide 9: Laboratory Wave-Meter (it is a very-accurate interferometer!)
Relative uncert. thanks to He-Ne/I2
x=1.55m x=1.5510-16m=0.15fm
x193THz x 19.3kHz 20kHz
Radiation under test (x) and a reference one (R) travel, in vacuum, along the same path counted as interference fringes on the two detectors PD1 e PD2

PD2 counts Nx fringes at x
PD1 counts (NR+) fringes at R
4L=Nxx=NRR
For R=633nm, L=30cm one gets NR1.9106 with x=1550nm one gets Nx7.7105
ur(R)=2.510-11
L
 /=10-10  ...
Michelson Wave-Meter 2 interferometers at x, R

## Slide 10: Wave-Meter for DWDM systems
Simultaneous meas. up to 256 channels
u()=0.3 pm (300fm!) and =10.1 pm
Meas. peak power and total power
Automatic meas. SNR and spacings 
Waterfall diagram of (t) and P(t)

## Slide 11: Spectrometer and monochromator
Tunable monochromator
=Kd
Kd=(d /d)
dispersive power
(spectral) resolving power r = | min /  | = | min /  | spectral res. ( angular res.) typ0.01nm
(10 pm >> 0.1-100 fm=RW-M)
n=n()

## Slide 12: OSA: functional schemes
traditional OSA
(monochromator with
rotating prism/grating)
narrow-band “field” OSA
shock-resistant
(fixed grating and moving detector or better fixed CCD array)

## Slide 13: Ultra-compact OSA (PC board)
optical fiber
lens
fixed
grating
CCD array
( Si, InGaS, IR-pushed-InGaAs )
Q, V, I  P

R
B
G
2048 pixel
(12.5m200m)
L=25.6mm
 res=0.3nm
 SPAN=600nm

## Slide 14: Laboratory OSA
Meas. of continuum spectrum (lines+ASE)
u()=10 pm e =2 pm
Meas. of power: peaks and total
Automatic meas. of SNR and  and P (markers)
Quantitative graphical representation

## Slide 15: Optical spectrum measurements for laser lines or D–WDM signals
Parameters to be measured (on the optical WDM signal):
channel power [dBm]
center wavelength [nm]
channel spacing [nm , GHz]
ASE noise floor [dBm/nm]
ratio S/N [dB/nm]
close channels crosstalk [dB]
global optical power [dBm]
Measurement parameters of the OSA:
center wavelength [nm] START [nm]
wavelength span [nm] STOP [nm]
resolution bandwidth (FWHM) [nm]
wavelength accuracy [nm]
saturation power [dBm]
"sensitivity" [dBm]
dynamic range [dB]
power accuracy [dB]
Optical Spectrum Analyzer (OSA)

## Slide 16: Optical Time Domain Reflectometry
Indirect measurement of local losses and distributed attenuation along an optical fiber cable.
Signal reading is done in reflection due to the backscattering.
In each “point” scattering is proportional to the available power level and depends inversely on the adopted wavelength.
Popt(t/2)
backscatter
Pbs(t)
 | S [dB/km] | RS [dB]
MM 850nm | 1.9 | -50
SM 1300nm | 0.35 | -63
SM 1550nm | 0.175 | -66
RS
spatial extension z of the “lit” zone providing for backscattering (with coeff. S per unit length)
z=vg
z
scattering coeff. per unit length .
causes...
vg/k con k=2/
×2
×5
×20
×2

## Slide 17: OTDR Pulse length and resolution
With a pulsed laser we inject short (ns) light pulses in the fiber (cable) and we observe the plot of “back-reflections” (backscattering) as a function of time: recorded time gives the distance z of the fiber length z
attenuation

z =  (c/n)
 = 15 ns
z=2 m
z=0.66 km
 = 33 s
< energia
> energia
z
z
Spatial resolution z is poorer when the pulse is larger but SNR gets better with larger/stronger pulses

## Slide 18: Information contained in the OTDR trace
“A picture is worth a thousand words”
Pb. saturation
and line-base recovering
t= =z / (c/n)

## Slide 19: Tricks in OTDR measurements
joints between different fibers
artifacts ‘ghosts’ due to multiple reflections

## Slide 20: Example of OTDR measurement
portable OTR for “in-field”
measurements
measurement panel of an OTDR

## Slide 21: Fiber Insertion Loss measurement
Results of different measurements
at 850 nm on a 520 m cable
Test | Loss , Std. Dev.
1-Cable Reference | 2.96 dB , 0.02 dB
2-Cable Reference | 2.66 dB , 0.20 dB
3-Cable Reference | 2.48 dB , 0.24 dB
OTDR (1/2 directions) | 2.05 dB / 1.91 dB
4 standardized methods exist

## Slide 22: PMD (Polarization Mode Dispersion) meas.
PMD: STATISTICAL parameter
Important Avg Value and St.Dev. of DGD (Differential Group Delay) We measure variance per unit length (in time or as a function of SOP or ): hence we get 2 in [ps2/km] and then PMD=DGD )
Typical values PMD 0.2 ps/(km)1/2
variability range 0.11 ps/(km)1/2
Opt. Commun. >10 Gb/s
Pb.: polariz. dispersion
meas. with multiple 
meas. with a single 
we change “all” possible polarization states (SOPs)
media e varianza dei ritardi
or Time
t
ny  nx
PAR.
SEQ.

## Slide 23
BER and eye-diagram measurements
For BER<10-12 long times are needed (>>20 min at 1 GSa/s)…
BER can be estimated from the opening (S/N ratio) of the eye-diagram
eye-diagram opening
time
jitter

## Slide 24: D-WDM System Monitoring
Wave–Meter
Power
Meter
OTDR
Notes: A conclusione,
possiamo dire che abbiamo descritto i principi fisici su cui si basa l’azione laser e i dispositivi che consento di costruire una sorgente laser in continua.
Quindi, abbiamo visto le principali proprietà dei fasci laser e le misure: di potenza ottica, con gli associati rivelatori, di profilo trasversale, della lunghezza d’onda e dello spettro ottico.
Sono poi stati descritti alcuni metodi e sistemi per la stabilizzazione, in ampiezza e frequenza, delle sorgenti laser .
Da ultimo, abbiamo riportato alcuni esempi di campioni di frequenza ottica.
Se dopo l’invenzione del laser e le sue prime realizzazioni si usava dire che: “*”,
oggi possiamo dire che gli strumenti laser trovano vasta applicazione in diversi campi della ricerca scientifica, nel settore tecnologico-produttivo e per le misure opto-elettroniche.

## Slide 25: Bibliography
S. Donati, Photodetectors: Devices, Circuits and Applications, Prentice Hall, New York, 2000
W. Demtröder, Laser Spectroscopy, Springer, Berlin, 1996
Application Note 1550-4, Optical Spectrum Analysis - Basics, Hewlett Packard, U.S.A., 1996
http:www.iec.org/tutorials/dwdm_test
Notes: Riportiamo, infine, i riferimenti ad alcuni testi scientifici che possono essere usati per approfondimenti sugli argomenti presentati.
E’ stato un privilegio e un piacere insegnare in questa Scuola.
Gazie per la vostra attenzione.